A full range calibration for CALIFA
Mrunmoy Jena
R3B Week
12.11.2024
Overview and structure of macro
Calibration in the gamma range
Extrapolation to proton range
Physics with R3B
CALIFA – Design
TUM Members:
Roman Gernhäuser,Philipp Klenze,Tobias Jenegger, Mrunmoy Jena

---
Introduction
511 keV
Proton range
22Na source
Gamma range
22Na source
Source peaks
1275 keV
22Na source hardly useful for performing calibration in the proton range !

---
Introduction
511 keV
Proton range
22Na + pulser
Gamma range
22Na + pulser
22Na peaks
1275 keV
P1
P2
P3
P1
P5
P6
P2
P4
P3
Using pulsers is essential for calibration in proton range

---
Introduction
HVAPD
Integrator stage
CAPD
HVAPD
L(t)
Rf
RAPD
Cf
Gain stage
Vout
CP
R1
Pulser
L(t)≡ f(Nγ ⋅ cTransp. ⋅ cAbsorp. ⋅ ϵAPD ⋅ gAPD )
R3
R2
Need to calibrate:
S
\mathbf{\frac{\int L(t) dt}{V_{out}}}

---
Calibration overview
40 min
4 hrs
2 min
unpacked.root
(60 GB)
QC-CALIFA.root
(30 MB)
spectra.root
(60 MB)
data.lmd
(90 GB)
Key advantages:
Unpacking, writing spectra done only once !
Calibration takes just 2 min.

---
Calibration overview
Calibration macro
Unpacking
Run bash script
(bash_pulser_calibration.sh)
Unpack
using
UCESB
reader
unpacked_gamma.root
unpacked_proton.root
data.lmd
read all events
and fill
Pulser region
Source region
spectra_gamma.root
spectra_proton.root
(5088 1D histograms)
Mean, Sigma, Events
for each histogram
Gauss fit each peak
find peaks
Peak parameters:
number and region for
source and pulser peaks
Fit parameters:
Expected width

---
Calibration macro
Results
Fit function:
y = c0 + mx
Fit function:
y = c + mx
Write to .par file
For single peak fit/ gain match
For multiple peak fit
QC-CALIFA.root
With 2D histograms of calibrated energy, slope, offset, range factor …
calib.par
With calibration parameters : slope and offset
Gamma gain fit
Source
energies
Slopegamma, offsetgamma
Range factor
Gamma pulser peak positions
Slopeproton, offsetproton
Mean value (σ) of
 all peaks
Proton gain fit
Write to .par file
Fit function:
y = c + mx
40 min
4 hrs
2 min
data.lmd
(90 GB)
unpacked.root
(60 GB)
QC-CALIFA.root
(30 MB)
spectra.root
(60 MB)

---
Range factor
3/45 pC
3/45 pC
Back barrel
(single range)
3/30 pC
CEPA
(dual
range)
3/30 pC
3/30 pC
Barrel
(single range)
Tolerance (σ/μ)
= 6 ‰
iPhos
(dual range)
Quality of resistors
in the gain stage
R=Eγ (uncalib.)/Ep (uncalib.) ≈ 11
Serves as a translation factor between
gamma range and proton range

---
Gamma range: uncalibrated spectra
22Na – 1275 keV peak
CEPA
iPhos
Barrel
Back barrel
Gain matched channels
22Na – 1275 keV peak
σ/μ = 4%

---
Gamma range: calibrated spectra
CEPA
22Na – 1275 keV peak
Back barrel
iPhos
Barrel
Applying a linear fit: y=mγ x+cγ
Software works !
Now we also have pulser peaks in keV

---
Proton range: uncalibrated spectra
Barrel
iPhos
CEPA
Back barrel
There are (channel to channel) variations in the electronics

---
Calibration in proton range
Back
barrel
CEPA
iPhos
Barrel
σ/μ = 3%
Energies for first 3 pulsers obtained from gamma calibration
Then doing a linear fit : y=mpx+cp ⇒ All pulser energies in keV
Variation in pulsers include all uncertainities (depending on L(t), R and Cpulser)

---
Circuit diagram
HVAPD
Integrator stage
CAPD
HVAPD
L(t)
Rf
RAPD
Cf
Gain stage
Vout
CP
R1
Pulser
L(t)≡ f(Nγ ⋅ cTransp. ⋅ cAbsorp. ⋅ ϵAPD ⋅ gAPD )
R3
R2
S

---
Pulser capacitance
σ/μ = 3%
Variation in E/V for the pulsers again include all uncertainities (L(t), R and Cpulser)

---
Proton range: calibrated spectra
CEPA
iPhos
Barrel
Back barrel
σ/μ = 0.4 ‰
Software works !

---
QC Plots
Back
barrel
iPhos
Barrel
CEPA
iPhos
CEPA
Barrel
Barrel
PC 0
PC 1
DAQ Plot (Hardware level)
Quality Assessment Plot
WIX
MES

---
QC Plots
Back
barrel
iPhos
CEPA
Barrel
Barrel
511 keV
P1
1275 keV
Spectrum for bad channel
(incorrect gain matching)
PC 0
PC 1
DAQ Plot
WIX
MES

---
Thank You !
CALIFA @ Technical University of Munich
Roman Gernhäuser,Philipp Klenze,Tobias Jenegger, Mrunmoy Jena

---
Extras: Bad channel
Crystal ID: 5057, in CEPA

---
Extras: Range factor vs bin number